## Matrix
[7 empty] | 202 | 203 | 202 | 202 | 218 | 218 | 210 | 210 | 219 | 218 | 218 | 218 | 213 | 219 | 219 | 219 | 219 | 219 | 219 | 217 | 219 | 213 | 213 | 218 | 219 | 219 | 218 | 208 | 208 | 207 | 207 | 207 | 207 | 207 | 201 | 200 | 199 | 199 | 202 | 202 | 197 | 200 | 203 | 203 | 203 | 200 | 200 | 201 | 201 | 201 | 204 | 204 | 211 | 201 | 198 | 198 | 205 | 204 | 204 | 204 | 201 | 204 | 204 | 204 | 185 | 186 | 88 | 88 | 126 | 126 | 124 | 94 | 11 | 11 | 10 | 28 | 30 | 25 | 25 | 19 | 27 | 27 | 29 | 23 | 12 | 12 | 56 | 88 | 37 | 36 | 13 | 20 | 22 | 12 | 11 | 39 | 30 | 31 | 28 | 22 | 21 | 82 | 89 | 104 | 71 | 75 | 28 | 70 | 70 | 69 | 67 | 76 | 75 | 60 | 17 | 20 | 29 | 29 | 33 | 33 | 9 | 27 | 36 | 20 | 38 | 38 | 32 | 61 | 50 | 50 | 19 | 35 | 34 | 23 | 23 | 9 | 22 | 19 | 32 | 29 | 24
NEW !!!
Cell Count |  |  |  |  | Modified Information |  | System | System | System | System | Phased Array | Phased Array | Phased Array | Phased Array | Linear | Linear | Linear | Linear | Linear | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Curved | Curved | Curved | Curved | Curved | Curved | Linear | Linear | Linear | Curved | Curved | Curved | Curved | Curved | Linear | Linear | Linear | Endocavity | Endocavity | Endocavity | Volume | Volume | Volume | Volume | Volume | Volume | Phased Array | Linear | Linear | Linear | Linear | Linear | Linear | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Curved | Curved | Curved | Curved | Linear | Curved | Linear | Linear | Curved | Curved | Curved | Linear | Linear Button | Linear Button | Biplane | Biplane | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Endocavity | Endocavity | Endocavity | Endocavity | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | CW | CW | CW | CW | CW | TEE | TEE | TEE
16809 | Classification | Name | Type | Disinfectanct Level | Manufacturer (Country) | Active Ingredient | Body Surface | Control Panel | Control Panel HERA/RS : '23.Feb~ V8/V7 : '23.Jul~ HS40/HS50/HS60 : '24.Dec~ Z20/R20/V6/V5 | Monitor | PA1-5A | PA1-5AE | PA2-9S | PA3-15S | LA2-14A | LA3-16A | LA3-16AD | LF5-13 | LA2-16S | EA2-11AR | EA2-11AV | EA2-11ARD | EA2-11ARE | EA2-11AVD | EA2-11AVE | EA2-11ARn | miniER7 | CM2-11 | CA2-13M | CA3-10A | CA2-8A | CA2-8AD | CF2-8 | LA3-14A | LA3-14AD | L3-12A | CA2-9A | CA2-9AD | CA1-7A | CA1-7AD | CM1-8 | LM2-18 | LM2-18D | LM3-27 | EV2-10A | EV2-12 | EV3-15 | VN4-8 | CV1-8A | CV1-8AD | CV1-8AE | CMV1-10 | CMV1-10Z | PN2-4 | L5-13/50 | L5-12/50 | LN5-12/40 | LN5-12 | LA2-9S | L3-8 | EA2-12SR | EA2-12SV | ER4-9 | EV4-9 | EVN4-9 | EC4-9 | CN4-9 | C2-8 | C2-8ED | C2-5 | L3-22 | CA4-10M | LA4-18A | LA4-18AD | CA1-7S | CA1-7SD | CA1-7Sn | LA3-22AI | LA2-16Sc | LA3-16ADc | BCC2-13 | BCL2-14 | 3DC2-6 | 3D2-6 | 3D2-6ET | V2-6 | 3D4-8 | 3D4-8ET | V4-8 | VE4-8 | 3D4-9 | 3D4-9ES | V5-9 | EV3-10B | LV3-14A | V6-12 | C1-4EC | CS1-4 | SC1-6 | C2-4/20 | C2-4ES | CA2-6BM | C2-6 | C2-6IC | CN2-8 | C3-7EP | C4-9/10ED | C5-8 | CF4-9 | EC4-9IS | VR5-9 | EA2-11B | E3-12A | LA2-9A | L4-7 | LA3-16AI | LS6-15 | LA4-18B | LA4-18BD | LM4-15B | LF5-12 | LS5-13 | L5-13 | L5-13IS | L7-16 | L7-16IS | LA2-9R Vet | P1-4 | PM1-6A | P2-4AH | P2-4BA | PE2-4 | SP3-8 | PA3-8B | P3-8 | P3-8CA | PA3-9B | PA4-12B | P4-12 | DP2B | CW2.0 | CW 4.0 | CW6.0 | DP8B | TA2-9 | MMPT3-7 | MPT3-7
67 | Disinfectant | 3M SURFACE DISINFECTANT CLEANER WIPES | Wipes | LLD | 3M (Canada) | Quaternary Ammonium(N-Alkyl), Alcohol | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [37 empty] | ●
64 | Disinfectant | Accel INTERVention Wipes | Wipes | ILD | Diversey (Canada) | Hydrogen peroxide | ● | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ●
4 | Disinfectant | Accel Prevention Wipes | Wipes | ILD | Diversey (Canada) | Hydrogen peroxide | X | ● | ● | ●
67 | Disinfectant | Acrylan | Liquid | LLD | Antiseptica (Germany) | Quaternary Ammonium(N-Alkyl) |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [10 empty] | ○ | [15 empty] | ○ | ○ | ○
4 | Disinfectant | ALI-FLEX RTU | Liquid | LLD | Lalema (Canada) | Sodium Hypochlorite | X | X | ● | ●
2 | Cleaner | Alkazyme | Liquid | Cleaner | Alkapharm (France) | Quaternary Ammonium(N-Alkyl), Proteolytic Enzymes | [111 empty] | ○ | ○
42 | Disinfectant | Amity 3 step solution | Wipes | HLD | Amity (U.K) | Bis(3-aminoproply) dodecylamine, Didecyldimethyl Ammonium Chloride | ● | X | ● | ● | [23 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [39 empty] | ●
73 | Cleaner | Amity 3 step solution: Detergent | Wipes | Cleaner | Amity (U.K) | Alkyl ethoxylate | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [31 empty] | ●
67 | Disinfectant | Amity 3 step solution: HLD | Wipes | HLD | Amity (U.K) | Bis(3-aminoproply) dodecylamine, Didecyldimethyl Ammonium Chloride | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [37 empty] | ●
73 | Cleaner | Amity 3 step solution: Rinse | Wipes | Cleaner | Amity (U.K) | Deionised water | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [31 empty] | ●
50 | Disinfectant | Anios Clean Excel D | Liquid | LLD | Anios (France) | Alcohols, ethoxylated | X | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [17 empty] | ● | ● | [5 empty] | ●
67 | Disinfectant | Anios Excel Surface Premium | Liquid | LLD | Anios (France) | Quaternary Ammonium(N-Alkyl) | X | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○
4 | Disinfectant | ANIOSURF Premium NPC | Liquid | LLD | Anios (France) | Quaternary Ammonium(N-Alkyl) | ● | ● | ● | ●
41 | Cleaner | Aniosyme DD1 | Liquid | Cleaner | Anios (France) | Quaternary Ammonium(N-Alkyl), Proteolytic Enzymes |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [17 empty] | ○ | [11 empty] | ○ | ○ | [18 empty] | ○ | ○ |  | ○ | [13 empty] | ○ | ○ |  | ○ | ○
50 | Disinfectant | Aniosyme X3 | Liquid | LLD | Anios (France) | Didecylmethylpoly(oxethyl) Ammonium Propionate | X | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [17 empty] | ● | ● | [5 empty] | ●
66 | Disinfectant | Aniosyme XL3 | Liquid | LLD | Anios (France) | sulfamic acid, alcohols | ● | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ●
97 | Disinfectant | Anioxyde 1000 LD | Liquid | HLD | Anios (France) | Peracetic Acid | X | X | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | X | X | X | X | ○ | X | X | X | X | X | X | X | X | X | X | X | ● | ● | ● | ● | ○ | ○ | ○ | X | ● | ● | ● | ● | ○ | [9 empty] | ○ | ○ | ○ | ○ | [5 empty] | ○ | X | X |  |  |  | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | [10 empty] | ●
2 | Disinfection System | Antigermix E1 | UV | HLD | Germitec (France) | UV-C | [138 empty] | ● | ●
66 | Disinfection System | Antigermix S1 | UV | HLD | Germitec (France) | UV-C |  |  |  |  | ★ | ★ |  |  | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  |  |  |  | ★ | ★ |  |  | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ | ★ | ★ | ★ | [15 empty] | ★ | [14 empty] | ★ | ★ |  | ★ |  | ★ |  |  | ★ | ★ | [10 empty] | ★ |  |  | ★ | ★ |  | ★ | ★ | ★ |  | ★ | ★
69 | Disinfectant | Antigone wipes | Wipes | LLD | Dan Mor Natural Products and Chemicals Ltd (Israel) | Alcohol, Quaternary Ammonium(N-Alkyl) | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ●
63 | Disinfection System | AquaCide® / Ethos Automated Ultrasound Probe Cleaner Disinfector | Liquid | HLD | CS Medical (USA) | peracetic acid |  |  |  |  | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ | ★ | ★ | ★ | ● | ● | ● | ● | ● | ● | ● | ★ | ★ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ★ |  |  | ★ | ★ | ★ | ●
138 | Gel | Aquasonic 100 | Gel | Gel | Parker Laboratories, Inc. (USA) | NA | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ●
79 | Gel | Aquasonic Clear Ultrasound Gel | Gel | Gel | Parker Laboratories, Inc. (USA) | NA | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● |  |  |  |  | ● | [11 empty] | ● | [11 empty] | ● |  | ● | ● |  |  |  |  | ● | ● | [18 empty] | ● | ● | ●
35 | Disinfectant | Asepti-Wipes II | Wipes | LLD | Ecolab (U.K) | NA: Alcohol Quaternary Ammonium(N-Alkyl); ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ●: ●; ●: ●; ○: ●
4 | Disinfectant | Avert Sporicidal disinfectant cleaner | Liquid | ILD | Diversey (Canada) | NA: Sodium hypochlorite; X: ●; X: ●; ●: ●; ●: ●
4 | Cleaner | AZO detergent wipe | Wipes | Cleaner | Synergy Health Ltd (U.K) | NA: NA; X: X; X: ●; ●: ●; ●: ●
70 | Disinfectant | Azomax | Wipes | LLD | Synergy Health Ltd (U.K) | NA: Quaternary Ammonium(N-Alkyl), Alcohol; X: X; X: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ●: ●; ●: ●; ○: ●; ○: ●; ○: ○; ○: ●
84 | Disinfectant | Bacillol 30 Sensitive Foam | Spray | LLD | BODE Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [15 empty] | ● | ● | ● | [11 empty] | ● | ● | ○ |  |  |  | ● | ● |  |  | ● | ● | ● | [8 empty] | ●
84 | Disinfectant | Bacillol 30 Sensitive Green Tissue | Wipes | LLD | BODE Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [15 empty] | ● | ● | ● | [11 empty] | ● | ● | ○ |  |  |  | ● | ● |  |  | ● | ● | ● | [8 empty] | ●
84 | Disinfectant | Bacillol 30 Sensitive Tissue | Wipes | LLD | BODE Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [15 empty] | ● | ● | ● | [11 empty] | ● | ● | ○ |  |  |  | ● | ● |  |  | ● | ● | ● | [8 empty] | ●
81 | Disinfectant | Bacillol 30 Tissue | Wipes | LLD | BODE Chemie GmbH (Germany) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [30 empty] | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | [9 empty] | ●
71 | Disinfectant | Bacillol AF Tissues | Wipes | LLD | BODE Chemie GmbH (Germany) | NA: Alcohol; X: X; X: X; ●: ●; ●: X; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●
61 | Disinfectant | Bacillol Zero Tissues | Wipes | LLD | BODE Chemie GmbH (Germany) | NA: Organic Acids; X: X; X: X; ●: ●; ●: X; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●
68 | Disinfectant | Biguacid S | Liquid | LLD | Antiseptica (Germany) | NA: Quaternary Ammonium(N-Alkyl), Alcohol; X: X; X: X; ●: X; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ●: ○; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○
74 | Cleaner | BIOSPOT* | Tablet | Cleaner | HYDRACHEM (U.K) | Sodium Dichloroisocyanurate | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [39 empty] | ● | ●
2 | Disinfectant | Bleach 5.25% | Liquid | LLD | Worldwide Production (Nation_wide) | NA: Sodium Hypochlorite; ○: ○ | ○
72 | Disinfectant | CaviCide | Spray | LLD | Metrex (USA) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [20 empty] | ○ | ○ | [15 empty] | ○ | ○ | ○ | [7 empty] | ○
80 | Disinfectant | CaviWipes | Wipes | LLD | Metrex (USA) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | ● | ● | X | X | X | X | X | X | X | X | X | X | X | ● | ● |  |  | ● | ● | ● |  |  |  |  | ● |  |  |  | ● | [6 empty] | ● | ● | [13 empty] | ● | ● |  | ● | ● |  | ● | ● |  |  |  |  | ●
71 | Disinfectant | CaviWipes 1 | Wipes | LLD | Metrex (USA) | NA: alcohol, Quaternary Ammonium(N-Alkyl); X: X; X: X; ●: X; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ● | [7 empty] | ● | ● | ●
5 | Disinfectant | CaviWipes AF | Wipes | LLD | Metrex (USA) | NA: Quaternary Ammonium(N-Alkyl); ○: ●; ○: ●; ○: ●; ○: ●; ○: ●
17 | Disinfectant | Chlorine 1000PPM | Liquid | HLD | Worldwide Production (Nation_wide) | NA: Chlorine compound; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○
17 | Disinfectant | Chlorine Dioxide Solution | Liquid | HLD | vitalundfitmit100 Gmbh (Germany) | NA: Chlorine dioxide; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○
120 | Disinfectant | Cidex | Liquid | HLD | Advanced Sterilization Products (USA) | Glutaraldehyde | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(50) |  |  | ○ | ○(50) | ○ | ○ |  | ○ | ○ |  |  | ○ |  |  | ○ |  | ○ | ○(50) | ○(50) | ○ | ○(50) | ○(50) | ● | ○(50) | ○(50) | ○ | ○ | ● | ● | ○(50) | ○(50) | ○ | [5 empty] | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○(50) |  |  | ◆ | ◆ | ◆
129 | Disinfectant | Cidex OPA | Liquid | HLD | Advanced Sterilization Products (USA) | Ortho-phthalaldehyde | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ◆ | ◆ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ◆ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ◆ | ◆ | ◆ | ◆ | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(15) | ○ | ○ | ○(15) | ◆ | ○(15) | ○(15) | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(15) | ○(15) | ◆ | ○(15) | ◆ | ◆ | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○(15) | ○(15) | ○ | ○(15) | ○ | ◆ | ◆ | ◆
123 | Disinfectant | Cidex Plus | Liquid | HLD | Advanced Sterilization Products (USA) | Glutaraldehyde | [68 empty] | ○ | ○ | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○(25) | ○(25) |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(25) | ○(25) | ○ | ○(25) | ○(25) | ○ | ○(25) | ○(25) | ○ | ○ |  |  | ○(25) | ○(25) | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○(25) | ○(25) | ○ | ○(25) | ○ | bar: X to ●
91 | Cleaner | Cidezyme | Liquid | Cleaner | Advanced Sterilization Products (USA) | Proteolytic Enzymes | [68 empty] | ○ | ○ | ○ | [15 empty] | ● | [8 empty] | ○ |  |  | ○ |  | X | ● | ● | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | [8 empty] | ○ |  | ○ | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆ | bar: ● to ●
68 | Cleaner | CIDEZYME ECO | Liquid | Cleaner | Advanced Sterilization Products (USA) | Subtilisin (Protease) | [66 empty] | ● | ● |  |  |  | ● | bar: X to ●
72 | Cleaner | CIDEZYME Enzymatic Detergent | Liquid | Cleaner | Advanced Sterilization Products (USA) | di-Sodium tetraborate decahydrate | [68 empty] | ● | ● | ● | ● | [39 empty] | ● | ● | bar: X to ●
72 | Cleaner | CIDEZYME XTRA Multi-Enzymatic Detergent | Liquid | Cleaner | Advanced Sterilization Products (USA) | di-Sodium tetraborate decahydrate | [66 empty] | ● | ● | ● | ● | ● | ● | bar: X to ●
93 | Disinfectant | Cleanisept Wipe forte | Wipes | LLD | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [68 empty] | ● | ● | ● | ○ | ● | ● | ● |  | ● | [9 empty] | ● | ● | [7 empty] | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | bar: X to ●
90 | Disinfectant | Cleanisept Wipe forte Maxi | Wipes | LLD | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [68 empty] | ● | ● | ● | ○ |  |  |  |  | ● | [9 empty] | ● | ● | [7 empty] | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | bar: X to ●
95 | Disinfectant | Cleanisept Wipe maxi | Wipes | LLD | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | ● |  |  |  |  | ● | [5 empty] | ● | ●(15) |  |  |  | ● | ●(15) | ●(15) | [5 empty] | ● | ● | ● |  |  |  | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | bar: X to ●
95 | Disinfectant | Cleanisept Wipes | Wipes | LLD | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | ● |  |  |  |  | ● | [5 empty] | ● | ●(15) |  |  |  | ● | ●(15) | ●(15) | [5 empty] | ● | ● | ● |  |  |  | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | bar: X to ●
68 | Gel | Clear Image | Gel | Gel | NEXT Medical Products (USA) | NA | [75 empty] | ● | [38 empty] | ○ | bar: X to ●
68 | Disinfectant | Clinell Alcohol Wipes | Wipes | LLD | GAMA Healthcare Ltd (U.K) | Isopropyl Alcohol 70% | [68 empty] | ● | ● | ● | ● | bar: X to ●
69 | Disinfectant | Clinell Biodegradable Surface Wipes | Wipes | LLD | GAMA Healthcare Ltd (U.K) | PTII & PTIV biocide | [66 empty] | ● | ● |  |  |  | ● | bar: X to ●
69 | Disinfectant | Clinell Clorox Wipes | Wipes | LLD | GAMA Healthcare Ltd (U.K) | Sodium Hypochlorite | [66 empty] | ● | ● |  |  |  | ● | bar: X to ●
67 | Cleaner | Clinell Detergent Wipes(Yellow) | Wipes | Cleaner | GAMA Healthcare Ltd (U.K) | Lauryl Polyglucose, Butoxydiglycol | [68 empty] | ● | ● | ● | bar: X to ●
42 | Disinfectant | Clinell Enhanced Pods | Tablet | HLD | GAMA Healthcare Ltd (U.K) | Advanced Peracetic Acid, APA | [71 empty] | ● | bar: X to ●; ● to ●; ● to ●
70 | Disinfectant | Clinell Hand & Surface Wipes 84 | Wipes | LLD | GAMA Healthcare Ltd (U.K) | Quaternary Ammonium Compounds | [68 empty] | ● | ● | ● | ● | bar: X to ●
73 | Disinfectant | Clinell Peracetic Acid Wipes | Wipes | HLD | GAMA Healthcare Ltd (U.K) | Citric acid | [66 empty] | ● | ● | ● | ● | ● | ● | [39 empty] | ● | bar: X to ●
129 | Disinfectant | Clinell Universal Santising wipes | Wipes | LLD | GAMA Healthcare Ltd (U.K) | Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | bar: X to ●
130 | Disinfectant | Clinell Universal Spray | Spray | LLD | GAMA Healthcare Ltd (U.K) | Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | bar: X to ●
130 | Disinfectant | Clinell Universal wipes | Wipes | LLD | GAMA Healthcare Ltd (U.K) | Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | bar: X to ●
69 | Disinfectant | Clorox Healthcare Versasure wipes | Wipes | LLD | Clorox (USA) | Quaternary Ammonium Compound | [66 empty] | ● | ● |  |  |  | ● | bar: X to ●
82 | Disinfectant | Clorox Healthcare® Bleach Germicidal Wipes | Wipes | LLD | Clorox (USA) | Sodium hypochlorite | [66 empty] | ● | ● | ● | ● | ● | ● |  |  |  | ● | [10 empty] | ● | ● | [13 empty] | ● | ● | ● |  |  |  | ● | ● | ● | ● | bar: X to ●
83 | Disinfectant | Clorox Healthcare® Hydrogen Peroxide Cleaner Disinfectant Wipes | Wipes | LLD | Clorox (USA) | Hydrogen Peroxide | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | X | X | X | X | ○ | X | X | X | X | X | X | X | X | X | X | X | ● | ● |  |  | ● | ● | ● | [17 empty] | ● | ● | [11 empty] | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ○ | [8 empty] | ●
4 | Disinfectant | Clorox regular (Clorox regular 1 : water 10) | Liquid | LLD | Clorox (USA) | Sodium Hypochlorite | bar: X to ●
4 | Cleaner | CN'F | Liquid | Cleaner | Neo Medical (Korea) | Alcohol | bar: X to ●
65 | Disinfection System | D45 | UV | HLD | UV Smart (Netherland) | UV-C | [66 empty] | ★ | ★ | ★ | ★ | ★ | ★ | bar: ● to ●; ● to ◆+; ● to ●
54 | Disinfectant | deconex 11 UNIVERSAL | Liquid | LLD | Borer Chemie AG (Switzerland) | Sodium hypochlorite | ● | X | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [12 empty] | ★ | ★ | ★ | ★ | ★ | ○ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ○ | ○
63 | Disinfectant | Descogen-I | Liquid | HLD | Antiseptica (Germany) | Potassium peroxymonosulfate, Citric Acid | [103 empty] | ○ | bar: ● to ●
63 | Disinfectant | Descoton Extra | Liquid | HLD | Dr.Schumacher GmbH (Germany) | Glutaraldehyde | [103 empty] | ○ | bar: ● to ●
71 | Disinfectant | Detro San AF Wipes | Wipes | LLD | Detrox (Turkiye) | Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | [40 empty] | ● | ● | bar: X to ●
66 | Disinfectant | Disinfectant 1 Wipes | Wipes | LLD | Ecolab (USA) | Benzyl Alcohol | bar: X to ●
4 | Disinfectant | Dismozon® plus | Powder | HLD | BODE Chemie GmbH (Germany) | Magnesium monoperoxyphthalate hexahydrate | bar: X to ●
81 | Disinfectant | Dispatch Hospital Cleaner Disinfectant Towels with Bleach | Wipes | LLD | Clorox (USA) | Sodium hypochlorite | [66 empty] | ● | ● | ● | ● | ● | ● | [15 empty] | ● | [13 empty] | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | bar: X to ●
4 | Disinfectant | Distel Disinfectant Wipes | Wipes | HLD | Tristel Solutions Ltd. (U.K) | Polymeric Biguanide Hydrochloride | bar: X to ●
1 | Disinfectant | Durvet Chlorhexidine | Liquid | Cleaner | Durvet (USA) | Chlorhexidine gluconate | [120 empty] | X
69 | Disinfectant | Ecobreeze Wipe | Wipes | LLD | World of Disinfectant (Russia) | Alcohol, Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | bar: X to ●
4 | Disinfectant | Ecolab Actichlor Tablets* | Tablet | LLD | Ecolab (U.K) | adipic acid, Sodium Carbonate | bar: X to ●
88 | Gel | EcoVue | Gel | Gel | HR® Pharmaceuticals, Inc. (USA) | NA | [66 empty] | ● | ● | ● | ● | ● |  |  |  |  | ● | [11 empty] | ● | ● | ● | [14 empty] | ● | ● | ● | ● | ● |  |  | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | bar: X to ●
81 | Disinfectant | ED Wipes | Wipes | LLD | MH Healthcare (Korea) | Quaternary Ammonium(N-Alkyl) | [66 empty] | ● | ● | ● | ● | ● | ● | [29 empty] | ● | ● | ● |  | ● |  | ● | ● |  |  | ● | ● | [14 empty] | ● | bar: X to ●
90 | Cleaner | EMpower | Liquid | Cleaner | Metrex (USA) | Propylene glycol | X | X | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [21 empty] | ○ |  |  |  |  | ○ |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ | ○ |  |  |  | ○ | ○ | [11 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆
74 | Cleaner | EmPower Foam | Liquid | Cleaner | Metrex (USA) | Polyethylene glycol | X | X | X | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ● |  |  |  |  | ○ | ○
71 | Cleaner | Endozime | Liquid | Cleaner | Ruhof (USA) | Proteolytic Enzymes | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [21 empty] | ○ | [15 empty] | ○ | ○ | ○ | ●
2 | Cleaner | Endozime Sponge | Sponge | Cleaner | Ruhof (USA) | Proteolytic enzymes | [138 empty] | ● |  | ◆
17 | Cleaner | Endozime Xtreme Power | Liquid | Cleaner | Ruhof (USA) | Proteolytic Enzymes | [103 empty] | ○ | bar: ● to ●; ● to ●
17 | Disinfectant | Envirocide | Spray | LLD | Metrex (USA) | Alcohol Quaternary Ammonium(N-Alkyl) | [103 empty] | ○ | bar: ● to ●; ● to ●
66 | Disinfectant | EnviroSafe Wipes | Wipes | LLD | 경창메딕스 (Korea) | Alkyl(ethylbenzyl)dimethylammonium chloride | [66 empty] | ● | ● | bar: X to ●
92 | Cleaner | Enzol® | Liquid | Cleaner | Advanced Sterilization Products (USA) | Proteolytic Enzymes |  |  |  |  | ◆ | ◆ | ● | ● | ◆ | ● | ● | ● | ● | ◆ | ◆ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ● | ● | ● | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ● | ● | ● | ◆ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ● | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ◆ | ● | ● | ● | [14 empty] | ● | [8 empty] | ○ |  |  | ○ |  |  | ● | ● | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | [8 empty] | ○ |  | ○ | ○ | ○ |  | ○ | ○ | ○ | ● | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆
64 | Disinfectant | Enzyme Pre-Cleaner Wipes | Wipes | LLD | HP&C (Korea) | Sodium Benzoate | X | X | ● | ● | bar: ◆ to ○
72 | Disinfectant | Enzysept PLUS Wipes | Wipes | LLD | HP&C (Korea) | Benzalkonium Chloride | ● | ● | ● | ● | [62 empty] | ● | ● | ● | ● | ● | ● | bar: ◆ to ○
4 | Disinfectant | Ethanol 80% | Liquid | LLD | Worldwide Production (Nation_wide) | Ethanol 80% | X | ● | ● | X
39 | Disinfectant | gigasept | Liquid | HLD | Schülke & Mayr GmbH (Germany) | Succindialdehyde, formaldehyde | [77 empty] | X | X |  | X | X | [5 empty] | ○(30) | bar: ◆ to ○; ●; ○ to ○
87 | Disinfectant | gigasept FF | Liquid | HLD | Schülke & Mayr GmbH (Germany) | Succindialdehyde | [77 empty] | X | X |  | X | X | ○ | ○(30) |  |  |  | ○(30) | ○(30) | ○(30) | ○ |  |  |  | ○ |  | ★ | ★ |  |  |  | ○(30) | ○(30) | ○ | ○(30) | ○(30) |  | ○(30) | ○(30) |  |  |  |  | ○(30) | [19 empty] | ○(30) | ○(30) | ○ | ○(30) | bar: ◆ to ○
55 | Disinfectant | gigasept pearls* | Powder | HLD | Schülke & Mayr GmbH (Germany) | Sodium percarbonate, Citric Acid | ● | X | ● | ● | [83 empty] | ○ | [13 empty] | ● | ● |  | ○ | ○ | bar: ◆ to ●; ●; ○ to ○
77 | Disinfectant | gigasept® AF | Liquid | LLD | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [79 empty] | ○ | [7 empty] | ○ | ○ | ○ | [11 empty] | ○(30) | ○(30) | ○ | ○ | ○ |  | ○(30) | ○(30) |  |  |  |  | ○(30) | [19 empty] | ○(30) | ○(30) |  | ○(30) | bar: ◆ to ○
63 | Disinfectant | gigasept® FF (new) | Liquid | HLD | Schülke & Mayr GmbH (Germany) | Succindialdehyde | [103 empty] | ○ | bar: ◆ to ○
64 | Disinfectant | gigasept® PAA concentrate | Liquid | HLD | Schülke & Mayr GmbH (Germany) | Hydrogen peroxide Acetic Acid | X | X | ● | ● | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★/◆ | ★/◆ | ★/◆+ | ★/◆+ | ★/◆+ | ★/◆+ | ★/◆+ | ★/◆+ | ★ | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ | ★ | ★ | ★/◆+ | ★/◆+ | ★/◆+ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★/◆+ | ★/◆+ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | [9 empty] | X | [11 empty] | ★ | [18 empty] | ★ | [31 empty] | ◆ | ◆ | ◆
69 | Disinfectant | gigasept® PowerSET3 | Wipes | HLD | Schülke & Mayr GmbH (Germany) | Hydrogen Peroxide / Acetic Acid / Peracetic Acid | X | X | ● | ● | [62 empty] | ● | ● | ● | ● | ● | bar: ◆ to ○
66 | Disinfectant | HEXANIOS G+R | Liquid | LLD | Anios (France) | Quaternary Ammonium(N-Alkyl) | ● | X | ● | ● | bar: ◆ to ○
67 | Disinfectant | HEXI HUB | Liquid | HLD | GAMA Healthcare Ltd (U.K) | Advanced Peracetic Acid | X | X | ● | ● | [67 empty] | ● | bar: ◆ to ○
10 | Disinfectant | Hibitane | Liquid | HLD | Schering-Plough Animal Health Ltd. (New Zealand) | Chlorhexidine gluconate | [101 empty] | ○(30) | ○(30) |  | ○(30) | ○(30) |  | ○(30) | ○(30) |  |  |  |  | ○(30) | [19 empty] | ○(30) | ○(30) |  | ○(30)
70 | Disinfectant | HIS Wipes | Wipes | LLD | Jun Medicare (Korea) | Choroalkylbenzyldimethylammonium | ● | X | ● | ● | [64 empty] | ● | ● | ● | ● | bar: ◆ to ○
85 | Disinfectant | HMC NF* | Powder | HLD | MDD COMPANY GmbH (Switzerland) | Sodium percarbonate | ● | X | ● | ● | [62 empty] | ● | ● | ● | ● | ● | ● | [14 empty] | ○ | ○ | [9 empty] | ● |  |  |  |  | ● | ○ | ○ | ○ |  | ● | ● |  |  | ● | ● | ● | [16 empty] | ● | bar: ◆ to ○
69 | Disinfectant | HORT Active Sporicidal Wipes | Wipes | LLD | Polysept (Russia) | Hydrogen Peroxide | ● | X | ● | ● | [62 empty] | ● | ● | ● | ● | ● | bar: ◆ to ○
2 | Cleaner | Hydrichloric acid fuming 37% | Liquid | Cleaner | Millipore / Merck (다국적기업) | Hydrochloric acid | [138 empty] | ● | ●
66 | Disinfection System | Hypernova Chronos | UV | HLD | Germitec (France) | UV-C | [66 empty] | ★ | ★ | ★ | ★ | ★ | [15 empty] | ★ | [14 empty] | ★ | ★ |  | ★ |  | ★ |  |  | ★ | ★ | [10 empty] | ★ |  |  | ★ | ★ |  | ★ | ★ | ★ |  | ★ | ★ | bar: ◆ to ◆; ◆ to ●; ◆ to ●; ○ to ○; ◆ to ◆+; ● to ●; ● to ○; ● to ○; ○
4 | Disinfectant | Incidin Foam | Spray | LLD | Ecolab (U.K) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ● | ●
4 | Cleaner | Incidin Liquid | Liquid | Cleaner | Ecolab (U.K) | Alcohol | X | X | ● | ●
75 | Disinfectant | Incidin oxyfoam | Foam | LLD | Ecolab (U.K) | Hydrogen Peroxide | X | ● | ● | ● | [62 empty] | ● | ● | ● | ● | ● | ● | [29 empty] | ● | ● | ● | bar: ◆ to ○
57 | Disinfectant | Incidin oxyfoam S | Foam | LLD | Ecolab (U.K) | Hydrogen Peroxide | X | ● | ● | ● | [97 empty] | ● | ● | ● | bar: ◆ to ○; ● to ○
80 | Disinfectant | Incidin oxywipe | Wipes | LLD | Ecolab (U.K) | Hydrogen Peroxide | X | ● | ● | ● | [64 empty] | ● | ● | ● | [17 empty] | ● | ● | [11 empty] | X | X | ● |  |  | ● |  |  | ● | ● | ● | ● | [9 empty] | ● | bar: ◆ to ○
76 | Disinfectant | Incidin oxywipe S | Wipes | LLD | Ecolab (U.K) | Hydrogen Peroxide | X | ● | ● | ● | [64 empty] | ● | ● | ● | [30 empty] | ● | ● | ● | [5 empty] | ● | ● | ● | ● | bar: ◆ to ○
54 | Disinfectant | Incidin OxyWipes NG | Wipes | LLD | Ecolab (U.K) | Hydrogen peroxide | X | X | ● | ● | [62 empty] | ● | ● | ● | ● | ● | [40 empty] | ● | ● | bar: ◆ to ●
68 | Cleaner | Instru-Zyme® Gel | Liquid | Cleaner | HARTMANN | Proteolytic Enzyme | X | X | ● | ● | [62 empty] | ● | ● | bar: ◆ to ○
72 | Disinfectant | ISO Wipe | Wipes | LLD | Halyard (Australia) | Isopropyl alcohol 70% | X | X | ● | ● | [62 empty] | ● | ● | ● | ● | ● | ● | [29 empty] | ● | ● | bar: ◆ to ○
111 | Cleaner | Isopropyl alcohol 70% | Liquid | Cleaner | Worldwide Production (Nation_wide) | Isopropyl alcohol 70% | X | X | ● | ● | [62 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ○ | [10 empty] | ○ |  |  |  |  | ○ | [5 empty] | ○ |  | X | ○(20) | ○(20) | ○ | ○ | ○ |  | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ● |  |  | ○(15) | ○(15) | ○ | ○ |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(15) | ○(15) |  | ○(15) |  | ● | ● | bar: ◆ to ○
89 | Cleaner | Klenzyme | Liquid | Cleaner | Steris (USA) | Proteolytic Enzymes | [76 empty] | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○(10) | ○(10) | [7 empty] | ○ |  |  | ○ | ○ | X |  |  | ○ | [5 empty] | ○ | ○ | ○ | ○ | [8 empty] | ○ |  | ○ | ○ | ○ |  | ○ | ○ | ○ | ● | ○ | ○ |  |  |  |  | ○ | bar: ◆ to ○
61 | Disinfectant | klorsept 17 | Tablet | HLD | Medentech(Ireland) | Troclosene Sodium | ● | ● | ● | ● | [64 empty] | ● | ● | ● | bar: ◆ to ●; ● to ○
4 | Cleaner | Kohrsolin FF | Liquid | Cleaner | HARTMANN | Enzyme Blend, Alcohol Ethoxylate | X | X | ● | ●
56 | Disinfection System | Lumicare ONE | UV | HLD | LUMICARE (Australia) | UV-C | [68 empty] | ★ | ★ |  | ★ | bar: ◆ to ◆; ◆ to ◆+; ● to ●; ● to ●; ◆ to ◆+; ● to ○; ● to ○
4 | Disinfectant | Lysoform® | Liquid | HLD | Lysoform (Germany) | Formaldehyde | X | X | ● | ●
68 | Cleaner | Matrix Wipes | Wipes | Cleaner | Whiteley Corporation (Australia) | Unique Biofilm Removing Detergent | X | ● | ● | ● | [64 empty] | ● | ● | ● | ● | bar: ◆ to ○
100 | Cleaner | McKesson Alcohol Prep Pad Sterile | Wipes | Cleaner | McKesson Medical-Surgical Inc. (USA) | Alcohol | [68 empty] | ● | ● | ● | ○ |  |  |  |  | ○ | [10 empty] | ○ |  |  |  |  | ○ | [5 empty] | ○ |  | X | ○(20) | ○(20) | ○ | ○ | ○ |  | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○ |  |  | ○(15) | ○(15) | ○ | ○ |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(15) | ○(15) |  | ○(15) | bar: ◆ to ○
51 | Disinfectant | MEDI CLEAN | Liquid | LLD | UBM Cleaning Solution INC (Korea) | Quaternary Ammonium(N-Alkyl), Alcohol | [103 empty] | ○ | bar: ◆ to ○; ● to ○
66 | Disinfectant | Medi-Fect | Liquid | LLD | CS Medical (USA) | Decyldimethyloctylammonium chloride | X | X | ● | ● | bar: ◆ to ○
64 | Disinfectant | Medilox | Liquid | HLD | Soosan CMC (Korea) | Hypochlorous Acid | ● | X | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | ○ | ○ | ● | ● | ● | ● | ● | ●
63 | Disinfectant | Medipal Alcohol Disinfectant Wipes | Wipes | LLD | Pal International limited (U.K) | Alcohol | [103 empty] | ○ | bar: ◆ to ○
72 | Disinfectant | Medipal Chlorine Disinfectant Wipes | Wipes | LLD | Pal International limited (U.K) | Sodium chloride | ● | ● | ● | ● | [62 empty] | ● | ● | ● | ● | ● | ● | bar: ◆ to ○
62 | Disinfectant | Medipal Disinfectant Wipes | Wipes | LLD | Pal International limited (U.K) | Quaternary Ammonium(N-Alkyl), Alcohol | X | X | ● | ● | [91 empty] | ● | [5 empty] | ● | ● | ● |  |  | ○ |  |  |  |  | ● | ● | ○ | bar: ◆ to ○; ● to ○
69 | Disinfectant | Mediwipes DM | Wipes | LLD | Medilab Sp. Z.o.o (Poland) | Alcohol Quaternary Ammonium(N-Alkyl) | ● | X | ● | ● | [64 empty] | ● | ● | ● | bar: ◆ to ○
4 | Cleaner | Meliseptol rapid | Liquid | Cleaner | B. Braun Medical AG (Switzerland) | Alcohol | X | X | ● | ●
74 | Disinfectant | Meliseptol Wipes Sensitive | Wipes | LLD | B. Braun Medical AG (Switzerland) | Quaternary Ammonium(N-Alkyl) | X | ● | ● | ● | [64 empty] | ● | ● | ● | [30 empty] | ● | ● | ● | [7 empty] | ● | ● | bar: ◆ to ○
4 | Disinfectant | Melsept SF | Liquid | LLD | B. Braun Medical AG (Switzerland) | Quaternary Ammonium(N-Alkyl) | X | X | ● | ●
100 | Disinfectant | Metricide | Liquid | HLD | Metrex (USA) | Glutaraldehyde |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | X | ○ | ○ | X | X | X | X | X | X | X | X | X | X | X | ● | ● | [11 empty] | X | X |  | ○ | ○ | ○ | ○(50) |  |  | ○ | ○(50) | ○(50) | ○(50) |  | ○ |  |  |  |  | ○ | ○ |  |  |  | ○(50) | ○(50) |  | ○(50) | ○(50) |  | ○(50) | ○(50) |  |  | ○ | ○ | ○(50) | ○(50) | ○ | [5 empty] | ○ |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ |  | ○(50) |  | ◆ | ◆ | ◆
51 | Disinfectant | Metricide 28 | Liquid | HLD | Metrex (USA) | Glutaraldehyde | [82 empty] | ○ |  |  |  |  | ○ | [7 empty] | ○ | [8 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  |  |  | ○ | [8 empty] | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ◆ | ◆ | ◆ | bar: ◆ to ○; ●; ○ to ○
82 | Disinfectant | Metricide 30 | Liquid | HLD | Metrex (USA) | Glutaraldehyde | [66 empty] | ● | ● |  |  |  | ● | ● | ● | ● | [12 empty] | ○ | ○ | ○ | [13 empty] | ○ | [5 empty] | ○ | ○ | [5 empty] | ○ | ○ | [6 empty] | ○ | ○ | ○ | ○ | ○ | ○ | bar: ◆ to ○
94 | Disinfectant | Metricide OPA Plus | Liquid | HLD | Metrex (USA) | Ortho-phthalaldehyde | X | X | ● | ● | ● | ● | ● | ● | ◆ | ● | ● | ● | ● | ◆ | ◆ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ◆ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ◆ | ● | ● | ● | ● | ● | ● | ● | ○ | [9 empty] | ○ | ○ | [7 empty] | ○ | ○ | ○ | [5 empty] | ○ | ○ | ○ |  |  |  | X | X | [5 empty] | X | X | X | X | [18 empty] | ◆ | ◆ | ◆
84 | Disinfectant | Metricide Plus 30 | Liquid | HLD | Metrex (USA) | Glutaraldehyde | [86 empty] | ○ | ○ | [13 empty] | ● | ● | ● | ○ | ○ |  | ● | ● | ● | ● |  |  | ○ | [24 empty] | ◆ | ◆ | bar: X to ●; ◆ to ●; ●
29 | Disinfectant | MetriGuard | Spray | LLD | Metrex (USA) | Alcohol Quaternary Ammonium(N-Alkyl) | [103 empty] | ○ | bar: ● to ○; ○ to ○
83 | Cleaner | Metrisponge | Sponge | Cleaner | Metrex (USA) | Propylene glycol | [87 empty] | ○ |  |  |  |  | ○ |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ | ○ |  |  |  | ○ | ○ | [11 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | bar: ● to ●
92 | Cleaner | Metrizyme | Liquid | Cleaner | Metrex (USA) | Proteolytic Enzymes | [77 empty] | X | X |  | X | X | [5 empty] | ○ |  |  |  |  | ○ |  |  | ○ |  |  | ○ | ○ | X |  |  | ○ | ○ | ○ |  |  |  | ○ | ○ | [11 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆ | bar: ● to ●
71 | Disinfectant | Micro Kill Bleach Wipes | Wipes | LLD | Medline (USA) | Sodium Hypochlorite | [127 empty] | ● | ● | ● | bar: X to ○; ◆ to ●
81 | Disinfectant | Mikrobac Tissues | Wipes | LLD | BODE Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [87 empty] | ● | [14 empty] | ● | ● | ● | ● |  | ● | ● |  |  | ● | ● | [14 empty] | ● |  |  | ● | bar: X to ●; ◆ to ●
78 | Disinfectant | Mikrozid AF Wipes | Wipes | LLD | Schülke & Mayr GmbH (Germany) | Alcohol | [87 empty] | ● | [14 empty] | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | bar: X to ○; ● to ●
69 | Disinfectant | Mikrozid PAA Wipes | Wipes | LLD | Schülke & Mayr GmbH (Germany) | Peracetic Acid / Hydrogen Peroxide / Acetic Acid | bar: X to ○; ● to ●
32 | Disinfectant | Mikrozid sensitive liquid | Liquid | LLD | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [87 empty] | ○ | [15 empty] | ○ | ○ | ○ | bar: ● to ○; ○ to ○
85 | Disinfectant | Mikrozid sensitive Wipes | Wipes | LLD | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [87 empty] | ● | [13 empty] | ● | ● |  | ● | ● |  | ● | ● |  |  | ● | ● | [9 empty] | ● |  |  |  |  | ● | ● | ● | bar: X to ●
33 | Disinfectant | Mikrozid universal liquid | Liquid | LLD | Schülke & Mayr GmbH (Germany) | Alcohol | X: X; X: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ◆: ●; ◆+: ●; ◆+: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ●: ●; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○ | [26 empty] | ○
67 | Disinfectant | Mikrozid universal Wipes premium | Wipes | LLD | Schülke & Mayr GmbH (Germany) | Alcohol | bar: X to ○; ◆ to ●
87 | Disinfectant | Milton anti bacterial solution | Liquid | LLD | Milton (France) | Sodium Hypochlorite | [90 empty] | ○ |  |  |  | ○ |  | ○ | ○ | ○ |  | ○ |  |  | ○ | [5 empty] | ○ | ○ | [5 empty] | ○ | ○ | ○ | ○ |  | ○ |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | bar: X to ●
69 | Disinfectant | Mirodez Wipes | Wipes | LLD | World of Disinfectant (Russia) | Quaternary Ammonium(N-Alkyl) | bar: X to ○; ● to ●
55 | Gel | Natural Image | Gel | Gel | Sonotech Inc. (USA) | NA | [95 empty] | ○ |  |  |  | ○ | [23 empty] | ○ | [6 empty] | ● | [6 empty] | ○ | bar: ● to ●; ○; ● to ●; ●; ●
70 | Disinfectant | NDP SurfaPlus | Spray | LLD | Vesismin Health (Spain) | Isopropyl alcohol | [111 empty] | ● | ● | bar: X to ○; ◆ to ●
3 | Cleaner | Neodisher® MediClean forte | Wipes | Cleaner | Dr.Weigert (Germany) | Trisodium nitrilotriacetate | [138 empty] | ◆ | ◆ | ◆
52 | Disinfectant | Nilaqua Biodegradable Surface & Hand Wipes | Wipes | LLD | Waterless Limited (UK) | Benzalkonium Chloride | [111 empty] | ● | ● | [14 empty] | ● |  |  | ● | bar: X to ●; ○
1 | Disinfectant | Nolvasan®Solution Chlorhexidine diacetate | Liquid | Cleaner | Zoetis (USA) | Chlorhexidine diacetate | [120 empty] | ○
66 | Disinfectant | Nuclean | Liquid | LLD | Alkapharm (France) | Alcohol | [98 empty] | ○ |  | X | [8 empty] | X | X | bar: ● to ●
67 | Disinfectant | Omnicide - FG2 | Liquid | HLD | Coventry Chemicals Limited (U.K) | Glutaraldehyde |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | X | X | X | X | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [32 empty] | ○ |  | ○ |  |  | ○ | [5 empty] | ○ | ○
42 | Disinfectant | Omnicide 14 | Liquid | HLD | Metrex (USA) | Glutaraldehyde | [121 empty] | ○ |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | bar: ● to ○; ○; ● to ●
45 | Disinfectant | Omnicide28 | Liquid | HLD | Metrex (USA) | Glutaraldehyde | [95 empty] | ○ | ○ | ○ | [23 empty] | ○ |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | bar: ● to ○; ○; ● to ●
68 | Disinfectant | OPAL Disinfectant | Liquid | LLD | Whiteley Corporation (Australia) | O-phthaldehyde | bar: X to ○; ◆ to ●
57 | Gel | Opti Lube Gel | Gel | Gel | Optimum Medical (U.K) | NA | [99 empty] | ● |  | ● | ● | ● |  |  |  | ● | ● | bar: ● to ○; ○ to ●; ●
72 | Disinfectant | OptiGerm | Spray | LLD | Oleak (Brazil) | PHMB, Quarternary Ammonium | [107 empty] | ● | ● | bar: X to ○; ● to ●
72 | Disinfectant | Oxivir 1 | Wipes | ILD | Diversey (Canada) | Hydrogen Peroxide | bar: X to ●
94 | Disinfectant | Oxivir TB Wipes | Wipes | ILD | Diversey (Canada) | Hydrogen Peroxide | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | [10 empty] | ○ | ○ | ○ | ○ | [11 empty] | ○ | ○ | ○ |  | ○ | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  |  | ● | ● | ● | ●
5 | Disinfectant | Oxivir® Five 16 Concentrate | Liquid | LLD | Diversey (Canada) | Hydrogen Peroxide | [120 empty] | X | bar: X to ●
54 | Disinfectant | Peraflu D | Liquid | HLD | Velly Co.,Ltd (Korea) | Peracetic Acid | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ○ | [11 empty] | ● | ● | [45 empty] | ● | ● | [14 empty] | ● |  |  | ●
107 | Disinfectant | PeraSafe* | Powder | HLD | Lanxess (Germany) | Citric Acid | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | X | X | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | ● | ● | ● | ● | ○ | ○ | ○ | ● | ● | ● |  |  | ○ | X | X |  | X | X | ○ | ○(15) |  |  | ○(15) | ○(15) | ○(15) | ○(15) | [5 empty] | ○ | ○ | ○ |  |  |  | ○(15) | ○(15) | X | ○(15) | ○(15) |  | ○(15) | ○(15) | ○ | ○ |  |  | ○(15) |  | ○ |  |  | ○ | ○ | [13 empty] | ○(15) | ○(15) |  | ○(15) |  | ◆ | ◆ | ◆
4 | Disinfectant | perform®* | Powder | HLD | Schülke & Mayr GmbH (Germany) | Pentapotassium | bar: X to ●
62 | Disinfectant | PREempt HLD5 | Liquid | HLD | Virox Technologies Inc.(Canada) | 2% Accelerated Hydrogen Peroxide | bar: ● to ●
3 | Cleaner | Prolystica® 2X Concentrate Enzymatic Presoak & cleaner | Liquid | Cleaner | Steris (USA) | Ethanolamine Ethoxylated alcohol Polyalkylene glycol Glycerine | [138 empty] | ◆ | ◆ | ◆
85 | Disinfectant | Protex disinfectant spray | Spray | LLD | Parker Laboratories, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | X | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ● | ● | ● | [8 empty] | ○ |  |  | ○ |  |  |  | ○ | ○ | ○ | ○ |  | ○ | [9 empty] | ● | ● | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | [6 empty] | ○
74 | Disinfectant | Protex Pro Disinfection wipe | Wipes | LLD | Parker Laboratories, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | [86 empty] | ● | ● | [13 empty] | ● | ● | ● | ● | ● | [7 empty] | ● | bar: X to ●
82 | Disinfectant | PROTEX ULTRA DISINFECTANT WIPES | Wipes | LLD | Parker Laboratories, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | [101 empty] | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | bar: X to ●; ●
4 | Disinfectant | Pursept AF | Liquid | LLD | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | bar: X to ●
66 | Disinfectant | Quaternary Disinfectant Wipes | Wipes | LLD | Ecolab (USA) | Poly hydrochloride | bar: X to ●
1 | Disinfectant | Quatricide® PV Second Generation | Liquid | LLD | Pharmacal Research Laboratories (USA) | Quaternary Ammonium(N-Alkyl) | [120 empty] | X
57 | Disinfectant | Rapicide OPA/28 | Liquid | HLD | Medivators (USA) | Ortho-phthalaldehyde | [86 empty] | ○ | ○ | [11 empty] | X |  |  |  | ○ | ○ | ○ | [7 empty] | ○ | bar: ● to ○; ○ to ●
3 | Disinfectant | Rapicide PA HLD | Liquid | HLD | Medivators (USA) | Peracetic acid | [138 empty] | ◆ | ◆ | ◆
6 | Disinfectant | Rely+On Virkon* | Powder | HLD | LANXESS Corporation (Australia) | Potassium peroxymonosulfate | [90 empty] | ○ | [15 empty] | ○ | [9 empty] | ○ | ○ | ○ | ○
88 | Disinfectant | Revital-Ox® RESERT HLD | Liquid | HLD | Steris (USA) | Hydrogen Peroxide |  |  |  |  | ○ | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ○ | ○ | ○ | ● | ● | ● | ● | ● | [10 empty] | ○ | ○ | ○ | ○ | [13 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | [24 empty] | ◆ | ◆ | ◆
81 | Disinfectant | RHEOSEPT-WD plus Wipes | Wipes | LLD | Wachendorff-Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | [87 empty] | ● | ● | ● | [11 empty] | ● | ● |  |  |  |  | ● | ● |  |  | ● | ● | [9 empty] | ● | bar: X to ●
72 | Disinfectant | Rudocid | Wipes | LLD | Medidez (Russia) | isopropyl alcohol, benzalkonium chloride | [122 empty] | ● | bar: X to ●
57 | Disinfectant | Rudocid®ENDO | Wipes | LLD | Medidez (Russia) | Alcohol Quaternary Ammonium(N-Alkyl) | bar: X to ○; ○ to ●; ◆ to ●
66 | Disinfectant | S-7XTRA | Wipes | ILD | Steri-7 (UK) | Multi-actives | bar: X to ●
87 | Cleaner | saFecide Medical wipes | Wipes | Cleaner | Cosell Care (Korea) | Alcohol | [87 empty] | ● | [13 empty] | ● | ● |  |  | ● |  | ● | ● | ● |  | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | bar: X to ●
78 | Disinfectant | saFecide Sensitive wipes | Wipes | LLD | Cosell Care (Korea) | Quaternary Ammonium(N-Alkyl) | [87 empty] | ● | [13 empty] | ● | ● | ● |  | ● |  |  |  | ● |  | ● | ● | ● | bar: X to ●; ◆ to ●
11 | Disinfectant | Salvanios pH10 | Liquid | LLD | Anios (France) | Guanidinium acetate Quaternary Ammonium(N-Alkyl) | [111 empty] | ○ | ○ | [8 empty] | ○ |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | bar: ●
62 | Disinfectant | Sani Cloth active | Wipes | LLD | Ecolab (U.K) | Quaternary Ammonium(N-Alkyl) | [87 empty] | ● | [15 empty] | ○ |  | ● |  | ● | ● | bar: X to ○; ○ to ●; ◆ to ●
117 | Cleaner | Sani-Cloth 70 | Wipes | Cleaner | PDI (U.K) | Alcohol | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● |  |  | ● | ● | ● | ○ | [5 empty] | ○ | ○ | ○ | ● | ● | ○ | ○ |  |  | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ |  | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ |  | ○
85 | Disinfectant | Sani-Cloth Active | Wipes | LLD | PDI (U.K) | Quaternary Ammonium(N-Alkyl) | [79 empty] | ● |  |  | ● |  |  |  | ● | ● | ●(15) | ●(15) | [11 empty] | ● | ● | ● | ● | ● | [5 empty] | ● | ● | ● | bar: X to ●
67 | Cleaner | Sani-Cloth Detergent Wipes | Wipes | Cleaner | PDI (U.K) | NA | [103 empty] | ● | bar: X to ●
113 | Disinfectant | Sani-Cloth HB | Wipes | LLD | PDI (U.K) | Quaternary Ammonium(N-Alkyl) Tetrasodium EDTA | X | ● | ● | ● | ◆ | ● | ● | ● | ◆ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ◆ | ● | ● | ● | [5 empty] | X | X |  | X | X | ● | ●(30) |  |  | ◆ | ● |  |  |  | ● |  |  |  | ● |  |  | ● | ● | X | ●(30) | ●(30) | ● | ● | ● |  | ●(30) | ●(30) | ● | ● | ● | ● | ●(30) | ●(30) | ● | [5 empty] | ● |  | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ●(30) | ●(30) |  | ●(30) | ●
84 | Disinfectant | Sani-Cloth Plus | Wipes | LLD | PDI (U.K) | Alcohol Quaternary Ammonium(N-Alkyl) | [86 empty] | ● | ● | [15 empty] | ● | ● | ● | [7 empty] | ● | [7 empty] | ● |  |  | ● | ● |  | ● | ● | ● |  | ● | ● | bar: X to ●; ◆ to ●; ●
67 | Disinfectant | Sani-Cloth Prime | Wipes | LLD | PDI (U.K) | Isopropyl alcohol | bar: X to ○; ◆ to ●
3 | Disinfectant | Sani-Cloth Universal wipes | Wipes | LLD | PDI (U.K) | Quaternary Ammonium | bar: X to ●
69 | Disinfectant | Sani-Cloth® AF Universal | Wipes | LLD | PDI (U.K) | 2-Aminoethanol, D-D-A-C, Potassium, Phosphoric Acid | [127 empty] | ● | bar: X to ○; ◆ to ●
97 | Disinfectant | Sani-Cloth® AF3 Germicidal Disposable Wipe | Wipes | LLD | PDI | Quaternary Ammonium(N-Alkyl) | [82 empty] | ● |  |  |  | ● | ● | [7 empty] | ● | ● | ● |  |  |  | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | bar: X to ●; ◆ to ●; ○
102 | Disinfectant | Sani-Cloth® Bleach Germicidal Disposable Wipe | Wipes | LLD | PDI (U.K) | Sodium Hypochlorite | [82 empty] | ● |  |  |  | ● | ● | [7 empty] | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | bar: X to ●; ○
126 | Gel | Scan | Gel | Gel | Parker Laboratories, Inc. (USA) | NA | X | X | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ● |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○
61 | Disinfectant | Scepto Shield-1W Wipes | Wipes | LLD | Sceptre Medical (India) | Quaternary Ammonium(N-Alkyl) | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [8 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | [40 empty] | ● | ●
67 | Disinfectant | SciCan Optim 33TB Surface Cleaner Disinfectant Solution | Liquid | LLD | SciCan (Canada) | Hydrogen peroxide | X | X | X | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [37 empty] | ○
67 | Disinfectant | SciCan Optim 33TB Surface Cleaner Disinfectant Wipes | Wipes | LLD | SciCan (Canada) | Hydrogen peroxide | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | [5 empty] | ●
66 | Disinfectant | SCOP'ANIOS RC | Liquid | HLD | Anios (France) | Hydrogen peroxide, Peracetic Acid | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | X | X | X | X | ○ | X | X | X | X | X | X | X | X | X | X | X | ○ | ○
41 | Disinfectant | Sekusept Aktiv* | Powder | HLD | Ecolab (U.K) | NA: Sodium percarbonate, Citric acid; X: X; X: X; ●: ●; ●: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: X; ○: X; ○: X; ○: X; ○: X; ○: ●; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X
4 | Disinfectant | Sekusept Extra | Liquid | HLD | Ecolab (U.K) | NA: Glutaraldehyde | [5 empty] | X | X |  | X | X
2 | Cleaner | Sekusept Plus | Liquid | Cleaner | Ecolab (U.K) | NA: Glucoprotamin | [14 empty] | ○(30) | ○(30)
74 | Disinfectant | Septalkan | Wipes | LLD | Alkapharm (France) | Quaternary Ammonium(N-Alkyl) | ● | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | [30 empty] | ● | ● | ● |  |  |  | ● | ●
40 | Disinfectant | Septiwipes | Wipes | LLD | Dr.Schumacher GmbH (Germany) | NA: Quaternary Ammonium(N-Alkyl); ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ● | [10 empty] | ● | ●(15) |  |  |  | ● | ●(15) | ●(15) | [25 empty] | ●
2 | Disinfectant | SEPTOL LINGETTES | Wipes | LLD | ACTEON (Italy) | NA: Quaternary Ammonium(N-Alkyl) | [39 empty] | X | X
4 | Cleaner | SideKick Disinfecting Wipes | Wipes | Cleaner | Stryker Medical (USA) | NA: Isopropyl alcohol; X: X; X: ●; ●: ●; ●: ●
23 | Gel | Skintact Ultrasonic Gel | Gel | Gel | Leonhard Lang GmbH (Austria) | bar: NA; ○ to ○
58 | Disinfectant | Solucide | Spray | LLD | CS Medical (USA) | Diethylene glycol monobutyl Ether | X | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [8 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ●
2 | Disinfectant | Soluscope PAA | Liquid | HLD | Ecolab (U.K) | NA: Hydrogen peroxide, Acetic Acid, Peracetic Acid | [66 empty] | ◆ | ◆
74 | Disinfectant | Sono Ultrasound Wipes | Wipes | LLD | Advanced Ultrasound Solutions Inc. (USA) | Quaternary Ammonium(N-Alkyl) | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● |  |  |  | ● | [26 empty] | ● | ● | [7 empty] | ● | ● | [14 empty] | ● |  |  | ●
38 | Gel | Sonogel | Gel | Gel | Sonogel Vertriebs GmbH (Germany) | [76 empty] | ● |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ |  | ○ | [9 empty] | ○ | ○ | ● | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | ○ | [18 empty] | ○ | ○ |  | ○ | bar: NA; ○ to ○
1 | Disinfectant | Sporicidin® disinfectant solution | Liquid | Cleaner | Contec, Inc. (USA) | NA: Phenol | [48 empty] | ○
89 | Disinfectant | Sporox | Liquid | HLD | DSHealthcare Inc. (USA) | Hydrogen Peroxide | X | X | ● | ● | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | X | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | X | [10 empty] | X | X | X |  | X | X |  |  | ○ | ○ | [6 empty] | ○ | ○ |  | ○ |  |  |  | ○ | X | [5 empty] | ○ |  |  | ○ | ○ | [5 empty] | ○ | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | [10 empty] | ○
90 | Disinfectant | SporoxⅡ | Liquid | HLD | DSHealthcare Inc. (USA) | Hydrogen Peroxide | X | X | ● | ● | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [10 empty] | X | X | X |  | X | X |  |  | ○ | ○ | [6 empty] | ○ | ○ |  | ○ |  |  |  | ○ | X | [5 empty] | ○ |  |  | ○ | ○ | [5 empty] | ○ | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ |  |  |  | ● | [6 empty] | ○
51 | Disinfectant | Steranios 2% | Liquid | HLD | Anios (France) | NA: Glutaraldehyde; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ●; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ●; ○: ● | [14 empty] | ○ | ○ | [16 empty] | ○ | ○ | [5 empty] | ○ | ○ | ○ | [10 empty] | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | [5 empty] | ◆ | ◆ | ◆
1 | Disinfectant | Steranios 2% ECS | Liquid | HLD | Anios (France) | NA: Glutaraldehyde | [68 empty] | ◆
1 | Disinfectant | Steranios 2% NG | Liquid | HLD | Anios (France) | NA: Glutaraldehyde | [68 empty] | ◆
15 | Disinfectant | Ster-Bac Blu | Liquid | LLD | Ecolab (U.K) | NA: Quaternary Ammonium(N-Alkyl) | [5 empty] | X | X |  | X | X | ○ | ○(50) | [7 empty] | ○ | [15 empty] | ○(50) | ○(50) |  |  |  |  | ○(50) | ○(50) | ○ | [17 empty] | ○(50) | ○(50) |  | ○(50)
69 | Disinfectant | Sterisept Wipes | Wipes | LLD | AS Chemi-Pharm (Estonia) | Quaternary Ammonium(N-Alkyl) | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ●
4 | Disinfectant | Super Sani-Cloth Plus Wipes | Wipes | LLD | PDI (U.K) | NA: Quaternary Ammonium(N-Alkyl); X: X; X: ●; ●: ●; ●: ●
93 | Disinfectant | Super Sani-Cloth® Germicidal Disposable Wipe | Wipes | LLD | PDI (U.K) | Alcohol Quaternary Ammonium(N-Alkyl) | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [14 empty] | ● | ● |  |  | ● | [10 empty] | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● |  |  | ● |  |  |  |  | ● | ● | ●
26 | Disinfectant | SURFA' SAFE | Liquid | LLD | Anios (France) | NA: Quaternary Ammonium(N-Alkyl); X: ●; X: ●; ●: ●; ●: ●; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○ | [14 empty] | ○ | ○ | [15 empty] | ○ | ○ | ○ | [7 empty] | ○
1 | Disinfectant | Surfaces Hautes | Spray | HLD | Anios (France) | NA: Glutaraldehyde | [65 empty] | ○
34 | Disinfectant | Surfanios Premium | Liquid | LLD | Anios (France) | NA: Quaternary Ammonium(N-Alkyl), Alcohol; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○; ○: ○
74 | Disinfectant | SURFA'SAFE PREMIUM | Spray | LLD | Anios (France) | Quaternary Ammonium(N-Alkyl) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● |  |  | ● | ● | ● | ● | [23 empty] | ○ | [5 empty] | ● | ● | ○
4 | Disinfectant | Tana Apesin spray | Spray | LLD | tana Chemie GmbH (Germany) | NA: Quaternary Ammonium(N-Alkyl); X: X; X: X; ●: ●; ●: ●
3 | Disinfection System | TD-5® (w/TD100) | Liquid | HLD | CS Medical (USA) | NA: Glutaraldehyde | [66 empty] | ◆ | ◆ | ◆
1 | Disinfection System | TD-8® (w/TD100) | Liquid | HLD | CS Medical (USA) | NA: Ortho-phthalaldehyde | [66 empty] | ●
2 | Cleaner | TEEZyme Sponge | Sponge | Cleaner | CS Medical (USA) | NA: Proteolytic enzymes | [66 empty] | ● |  | ◆
4 | Cleaner | terralin liquid | Liquid | Cleaner | Schülke & Mayr GmbH (Germany) | NA: Alcohol; X: X; X: ●; ●: ●; ●: ●
37 | Disinfectant | terralin PAA | Liquid | HLD | Schülke & Mayr GmbH (Germany) | NA: Hydrogen peroxide, Acetic Acid, Peracetic Acid; X: X; X: X; ●: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X; ○: X
107 | Cleaner | Transeptic® Cleaning Solution | Spray | Cleaner | Parker Laboratories, Inc. (USA) | Isopropyl alcohol 70% | X | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ |  |  |  |  | ○ | [9 empty] | ○ | ○ |  |  |  |  | ○ | [5 empty] | ○ |  |  | ○(20) | ○(20) | ○ | ○ | ○ |  | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○(15) |  |  | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(15) | ○(15) |  | ○(15)
101 | Disinfectant | Tristel Duo ULT | Foam | HLD | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | X | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ◆ | ◆ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ◆ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | X | ○ | X | X | X | ○ | ○ | ● | ● | ◆ | ◆ | ◆ | ◆ | ● | ○ | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ● | ★ | ★ | ★ | ◆+ | ○ | [7 empty] | ○ | ○ | ◆ | ◆ | [7 empty] | ○ | ○ | ○ | X |  |  |  | ● | ◆ | ◆ | ◆ | ◆ |  |  | ○ | ○ | ● | ● | ○ |  |  | ○ | ○ | ○ | ○ | [6 empty] | ○ | ●
80 | Disinfectant | Tristel Fuse for Stella | Liquid | HLD | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | [87 empty] | ★ | [5 empty] | ● | ● | [11 empty] | ★ | ● | ● | ● | ● | bar: X to ●
89 | Cleaner | Tristel Pre-clean wipe | Wipes | Cleaner | Tristel Solutions Ltd. (U.K) | Alcohol | [86 empty] | ○ | ○ | [13 empty] | ● | ● | ○ | ○ | ○ |  | ● | ● | ● | ● | ● | ● | ○ | [13 empty] | ● | [10 empty] | ◆ | ◆ | bar: X to ●
85 | Cleaner | Tristel Rinse Wipes | Wipes | Cleaner | Tristel Solutions Ltd. (U.K) | Deionised water | [86 empty] | ○ | ○ | [13 empty] | ● | ● | ● | ○ | ○ | ● |  |  |  |  | ● | ● | ○ | [24 empty] | ◆ | ◆ | bar: X to ●
81 | Disinfectant | Tristel Solo Wipes | Wipes | LLD | Tristel Solutions Ltd. (U.K) | Quaternary Ammonium(N-Alkyl) | [101 empty] | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | bar: X to ●
96 | Disinfectant | Tristel Sporicidal Wipe | Wipes | HLD | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  |  | ○ | [7 empty] | ○ | ○ |  | ○ | [7 empty] | ○ | ○ | ○ |  |  |  |  | ● | ○ | ○ |  |  | ● | ● | ○ | ○ | ● | ● |  |  |  | ○ | ○ | ○ | ○ | [6 empty] | ○ | ● | [10 empty] | ◆ | ◆
75 | Disinfectant | Tristel Trio Wipes System | Wipes | HLD | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | ● | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | [9 empty] | ● | [11 empty] | ○ | [15 empty] | ○ | ○ | [6 empty] | ● | ● | [25 empty] | ◆ | ◆ | ◆
101 | Disinfectant | Tristel ULT | Foam | HLD | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | X | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ◆ | ◆ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ◆ | ◆+ | ◆+ | ● | ● | ● | ● | ● | ● | X | X | X | X | X | ○ | ○ | ● | ● | ◆ | ◆ | ◆ | ◆ | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ● | ★ | ★ | ★ | ◆+ | ○ | [7 empty] | ○ | ○ | ◆ | ◆ | [7 empty] | ○ | ○ | ○ | X |  |  |  | ● | ◆ | ◆ | ◆ | ◆ |  |  | ○ | ○ | ● | ● | ○ |  |  | ○ | ○ | ○ | ○ | [6 empty] | ○ | ●
92 | Cleaner | trophon® companion cleaning wipes | Wipes | Cleaner | Nanosonics Limited (Australia) | Quaternary Ammonium(N-Alkyl) | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [15 empty] | ● | ● |  |  | ● | [9 empty] | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ●
94 | Disinfection System | trophon® EPR | Vapor | HLD | Nanosonics Limited (Australia) | Hydrogen Peroxide |  |  |  |  | ● | ● |  |  | ◆ | ● | ● | ● |  | ◆ | ◆ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ★ |  |  | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ★ | ★ | ★ | ◆ | ★ | ◆ | ★ | ★ | ★ | ★ | ★ | ★ |  |  |  | ● | ● | ● |  |  |  | ◆ | ◆ | ◆ | ◆ |  | ● | ● | ● | ● | ● | ○ | ○ | ◆ | ● | ● | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ◆ | ◆ | ◆ | ◆ | ○ | ○ |  | ○ | ○ |  |  |  | ○ | ○ |  |  |  | ○ | ○ |  | ◆ | ◆ | ◆ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ | ○ |  |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○
94 | Disinfection System | trophon®2 | Vapor | HLD | Nanosonics Limited (Australia) | Hydrogen Peroxide |  |  |  |  | ● | ● |  |  | ◆ | ● | ● | ● |  | ◆ | ◆ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ★ |  |  | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ★ | ★ | ★ | ◆ | ★ | ◆ | ★ | ★ | ★ | ★ | ★ | ★ |  |  |  | ● | ● | ● |  |  |  | ◆ | ◆ | ◆ | ◆ |  | ● | ● | ● | ● | ● | ○ | ○ | ◆ | ● | ● | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ◆ | ◆ | ◆ | ◆ | ○ | ○ |  | ○ | ○ |  |  |  | ○ | ○ |  |  |  | ○ | ○ |  | ◆ | ◆ | ◆ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ | ○ |  |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○
94 | Disinfection System | trophon®3 | Vapor | HLD | Nanosonics Limited (Australia) | Hydrogen Peroxide |  |  |  |  | ● | ● |  |  | ◆+ | ● | ● | ● |  | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ◆+ | ★ |  |  | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ★ | ★ | ★ | ◆+ | ★ | ◆+ | ★ | ★ | ★ | ★ | ★ | ★ |  |  |  | ● | ● | ● |  |  |  | ◆+ | ◆+ | ◆+ | ◆+ |  | ● | ● | ● | ● | ● | ○ | ○ | ◆+ | ● | ● | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ◆+ | ◆+ | ◆+ | ◆+ | ○ | ○ |  | ○ | ○ |  |  |  | ○ | ○ |  |  |  | ○ | ○ |  | ◆+ | ◆+ | ◆+ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ | ○ |  |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○
117 | Disinfectant | T-Spray | Spray | LLD | Pharmaceutical Innovations, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | [77 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○(20) | ○(20) | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | X | ○(20) | ○(20) | ○ | ○ | ○ |  | ○(20) | ○(20) | ○ | ○ |  |  | ○(20) | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(20) | ○(20) |  | ○(20) | bar: X to ●; ○ to ○; ○
117 | Disinfectant | T-SprayⅡ | Spray | LLD | Pharmaceutical Innovations, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | [77 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○(20) | ○(20) | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | X | ○(20) | ○(20) | ○ | ○ | ○ |  | ○(20) | ○(20) | ○ | ○ |  |  | ○(20) | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(20) | ○(20) |  | ○(20) | bar: X to ●; ○ to ○; ○
4 | Disinfectant | Tuffie large detergent wipes | Wipes | LLD | Vernacare (UK) | Ethylan CPG, Betain 810 | bar: X to ●
41 | Disinfectant | Tuffies 5 | Wipes | LLD | Vernacare (UK) | Chlorhexidine Digluconate | bar: X to ●; ● to ●; ● to ●
67 | Disinfectant | U Klean | Wipes | LLD | Caremate (Korea) | Quaternary Ammonium(N-Alkyl) | [111 empty] | ● | ● | [17 empty] | ● | bar: X to ●
84 | Disinfectant | UltrOx HLD | Liquid | HLD | CIVCO (USA) | Hydrogen Peroxide |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ○ | ○ | ○ | ● | [14 empty] | ○ | ○ | ○ | ○ | [13 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | [24 empty] | ● | ◆ | ◆
78 | Disinfectant | UMONIUM38® NEUTRALIS SPRAY | Spray | LLD | Laboratoire Huckert’s International (Belgium) | Quaternary Ammonium(N-Alkyl) | [80 empty] | ● | ● | ● | [13 empty] | ● | ● |  |  |  | ● | ● |  |  |  |  | ● | ● | bar: X to ●; ○ to ○
79 | Disinfectant | UMONIUM38® Probes | Wipes | LLD | Laboratoire Huckert’s International (Belgium) | Quaternary Ammonium(N-Alkyl) | [80 empty] | ● | ● | ● | [13 empty] | ● | ● |  |  |  | ● | ● | ● |  |  |  | ● | ● | bar: X to ●; ○ to ○
69 | Disinfectant | Uniwipe Midi Bio Clinical Disinfectant Wipes | Wipes | LLD | Uniwipe (UK) | Quaternary Ammonium Compounds | bar: X to ●; ●
69 | Disinfectant | Uniwipe Midi Clinical Wipes | Wipes | LLD | Uniwipe (UK) | Quaternary Ammonium Compounds | bar: X to ●; ●
44 | Disinfectant | Uniwipe Midi Hand & Surface ​Disinfectant Wipes | Wipes | LLD | Uniwipe (UK) | Quaternary Ammonium Compounds | bar: X to ●; ● to ●; ● to ●; ○ to ○
46 | Cleaner | Uniwipe multi purpose Wipes | Wipes | Cleaner | Uniwipe (UK) | NA | bar: X to ●; ○ to ●; ●
69 | Disinfectant | Uniwipe Multi-Surface Disinfection Wipes | Wipes | LLD | Uniwipe (UK) | Quaternary Ammonium Compounds | bar: X to ●; ●
79 | Disinfectant | Velox Wipes NA | Wipes | LLD | Medisept (Poland) | Quaternary Ammonium(N-Alkyl) | [101 empty] | ● | ● |  |  | ● |  | ● | ● |  |  | ● | ● | bar: X to ●
63 | Disinfectant | VERT 2 GO ED | Liquid | LLD | Wood Wyant (Canada) | Quaternary Ammonium(N-Alkyl) | [103 empty] | ○ | bar: ● to ●
3 | Disinfectant | Virex® II 256 | Liquid | LLD | Diversey (Canada) | Quaternary Ammonium(N-Alkyl) | bar: ● to ●
79 | Disinfectant | Virkon S* | Powder | HLD | LANXESS Corporation (Australia) | Potassium peroxymonosulfate | [91 empty] | ● | [9 empty] | ● | ● | ● |  |  |  | ● | ● | [11 empty] | ○ | bar: X to ●
30 | Sterilization system | V-PRO | Vapor | Sterilization | Steris (USA) | Hydrogen peroxide | ●: ★; ◆: ★; ◆: ★; ◆+: ★; ◆+: ★; ◆+: ★; ◆+: ★; ●: ★; ●: ★; ●: ★; ○: ★; ○: ★; ○: ★; ○: ★; ○: ★; ○: ★; ○: ★; ●: ★; ●: ★; ●: ★; ○: ★; ●: ★; ●: ★; ○: ★; ○: ★; ●: ★/◆ | [32 empty] | ★ | ◆ | [16 empty] | ★ |  |  |  | ★
71 | Gel | Vue Ultrasound Gel | Gel | Gel | Optimum Medical (U.K) | Triethanolamine | [111 empty] | ● | ● | [17 empty] | ● | bar: X to ●; ○ to ○; ◆+
75 | Disinfectant | V-Wipes | Wipes | LLD | Whiteley Corporation (Australia) | Alcohol | [86 empty] | ○ | ○ | [16 empty] | ○ | ○ | [5 empty] | ● | ● | ○ | [13 empty] | ● | ● | ● | ● | bar: X to ●
2 | Gel | Wavelength | Gel | Gel | National Therapy Products Inc. (USA) | NA | [114 empty] | ○ | bar: ◆+
114 | Disinfectant | Wavicide®-01 | Liquid | HLD | Medical Chemical Corp. (USA) | Glutaraldehyde | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [10 empty] | ○ | X | X |  | ○ | ○ | ○ | ○(50) |  |  | ○(50) | ○(50) | ○(50) | ○(50) |  | ○ |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○(50) | ○(50) | ○ | ○(50) | ○(50) |  | ○(50) | ○(50) | X | X |  |  | ○(50) | ○(50) | ○ | [5 empty] | ○ |  | ○ | ○ | ○ |  | ○ | ○ | ○ | ● | ○ | ○ | ○(50) | ○(50) |  | ○(50) | ○ | ● | ● | ◆
82 | Cleaner | Wet Wipe Ethanol 70% Disinfection | Wipes | Cleaner | Wetwipe (Denmark) | Ethanol 70% | [101 empty] | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | [14 empty] | ● | ● | ● | bar: X to ○
82 | Disinfectant | Wet Wipe Triamin Disinfection | Wipes | LLD | Wetwipe (Denmark) | Quat Ammonium | [101 empty] | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | [14 empty] | ● | ● | ● | bar: X to ○
79 | Disinfectant | Wip' Anios Excel | Wipes | LLD | Anios (France) | Quaternary Ammonium(N-Alkyl) | [86 empty] | ● | ● | [13 empty] | ● | ● | ● | ● | ● | [7 empty] | ● | bar: X to ○
71 | Disinfectant | Wip' Anios Spor'Active | Foam | HLD | Anios (France) | Peracetic Acid | [101 empty] | ● | ● | ● |  |  |  | ● | ● | bar: X to ●
4 | Disinfectant | Yuhanrox regular (Yuhanrox regular 1: water 10) | Liquid | LLD | Yuhan clorox (Korea) | Sodium Hypochlorite | bar: X to ●
65 | Disinfectant | Zytec Healthcare Disinfecting Wipes | Wipes | LLD | Empack Spraytech, Inc. (Canada) | Quaternary Ammonium Compounds | bar: ● to ●; ●
34 | Disinfectant | 닥터소독티슈 | Wipes | LLD | Carepharm (Korea) | Quaternary Ammonium(N-Alkyl) | bar: ● to ○; ○; ● to ●
29 | Disinfectant | 디터점스울트라 | Liquid | HLD | 디엠바이오 (Korea) | Sodium Hypochlorite | bar: X to ●; ● to ●
4 | Disinfectant | 绿先锋消毒剂 | Liquid | HLD | Techgreen (China) | Chlorine dioxide | bar: ○ to ○; ○
38 | Disinfectant | 소프라이트 죠킨II(ソフライト除菌Ⅱ) | Wipes | ILD | HAKUJUJI(Japan) | Ethanol, Isopropanol | bar: ● to ●; ● to ●; ○ to ○
38 | Disinfectant | 쇼도쿠슈퍼(ショードックスーパー) | Wipes | LLD | SARAYA CO., LTD.(Japan) | Ethanol, Citric Acid | bar: ● to ●; ● to ●; ○ to ○
40 | Disinfectant | 오피덱스액 | Liquid | HLD | (주)퍼슨헬스케어 (Korea) | Ortho-Phthalaldehyde, OPA | bar: ● to ●; X to ●
66 | Cleaner | 페라플루와입스 | Wipes | Cleaner | 밸리 (Korea) | Quaternary Ammonium(N-Alkyl) | bar: X to ●
41 | Disinfectant | 헥시크린에이액 5% | Liquid | HLD | 구미제약 (Korea) | Chlorhexidine gluconate | bar: X to ●; ● to ●; ● to ●
System
X = Not compatible(DO NOT USE)
● = Compatible
Transducer
X = Not compatible(DO NOT USE)
Blank = Untested.(DO NOT USE)
○ = Compatible (From Lens to Housing)
● = Compatible (From Lens upto 50cm away from the connector)
○(0*) = Must not be used longer than (n) minutes
★ = Staining may occur on surface of transducer; however, the acoustic performance and image quality are not affected.
◆ = Compatible and Validated for disinfection method
◆/★ = Compatible and Validated but staining may occur on the exterior; however, the acoustic performance and image quality are not affected.
◆+= Efficacy cross-functionally Compatible and Validated for disinfection method
User Notice
* All transducers are under tested in IPx 7 Criteria.
* Be sure to check the manufacturer's instructions for the dilution strength and method for using powders including Perasafe.
* V-PRO products include V-PRO1, V-PRO1 Plus, V-PRO Max, V-PRO Max2, V-PRO 60 sterilizer. - For guaranteed performance, this device has been tested and confirmed for up to 200 cycles.
* Be sure to follow the manufacturer's instructions before using V-pro sterilization.
* Be sure to check the manufacturer's instructions before using Antigermix disinfection system to prevent discoloration.
* Do not use ultrasound gels or couplant that contain any of the following ingredients: - Oil-based lotion type couplants, Mineral oil ,Olive oil, Lubricant oil, Paraffin-based grease, Dimethylsilicone, Alcohol such as methanol, - Ethanol or isopropanol, Acetone ,Dioctylphtalate, Aromatic substances, Iodine compounds, Aloe Vera
* Non-recommended Gel - POLYSONIC® ULTRASOUND LOTION